Cermet II I.S. Hygrometer

Users Guide

TABLE OF CONTENTS

SECTION							    			      PAGE

1	INTRODUCTION								5

1.1	General								5

1.2	CERMET II I.S. Hygrometer Overview				5

1.3	TRANSMET I.S. Overview						5

1.4	Outline of Complete System					6

1.5	Calibration								7

1.6	Manufacturing Quality						7

2	PREPARING YOUR CERMET II MONITOR FOR USE		7

2.1	Installing the Monitor						7

2.2	Electrical Connections						8

2.2.1	Electrical Power							9

2.2.2	Setpoint (Alarm) Relays						9

2.2.3	Analogue Output						10

2.2.4	Digital Interface							10

3	TRANSMET I.S.								11

3.1	TRANSMET I.S. Identification					11

3.2	Mechanical Installation						11

3.3	Operating Precautions						12

3.4	Sampling Hints							12

3.5	Response Characteristics						13

3.6	Which Gases to Measure						13

3.7	Maintaining the TRANSMET I.S.					13

4	TRANSMET I.S. TERMINAL BLOCK CONNECTION			14

5	TRANSMET I.S. CURRENT OUTPUT					15

6	CONNECTING TRANSMET I.S. TO CERMET II MONITOR		16

TABLE OF CONTENTS Cont.

SECTION							    			      PAGE

7	CERMET II MONITOR SETUP						17

7.1	Setup Security Feature						17

7.2	Selecting Engineering Units					18

7.3	Changing the Setpoint Values					18

7.4	Analogue Output Scaling						20

7.5	Display Brightness Adjustment					20

7.6	Digital Communications						21

8	PRESSURE COMPENSATION						23

8.1	Using a Pressure Transducer					23

8.2	Using a Fixed Pressure Input in Single Channel Mode		24

8.3	Pressure Transducer Connection					24

9	MONITOR MAINTENANCE						24

10	TECHNICAL SPECIFICATIONS						25

10.1	CERMET II Monitor							25

10.2	TRANSMET I.S.							26

11	INTRINSICALLY SAFE SYSTEM 2 DRAWING				27

FIGURES

1	Complete System								6

2	Panel Cut-out								8

3	Case Dimensions								8

4	Instrument Wiring Diagram							9

5	Terminal Block Connections						14

6	Current Output Switch Settings						15

7	Connecting TRANSMET I.S. to CERMET II Monitor			16

8	Location of the LOCKOUT Switches					17

TABLE OF CONTENTS Cont.

SECTION							    			      PAGE

APPENDIX

1	Register Settings Accessible by Digital Communication		28

2	SETUP Codes								30

CAL										30

CODE 1									31

CODE 3									31

CODE 4 – Channel 1								32

CODE 6 – Channel 2								32

CODE 7									33

Setpoint Control 1-6								33

3	Recycling									35

4	Hazardous Products							35

5	Certification									35

1.	INTRODUCTION

1.1	General

Michell Instruments manufactures a wide range of on-line and portable instruments for the determination of moisture content in air and gases. In addition, the company manufactures humidity calibration systems for industrial and scientific users.

CERMET II I.S. utilises the TRANSMET I.S. intrinsically safe dewpoint transmitter as the sensor coupled with the CERMET II advanced monitor as the display unit.

1.2	CERMET II I.S. Hygrometer Overview

The CERMET II I.S. Hygrometer is a continuous, on-line instrument for the measurement of absolute moisture content in gas. It is designed to fulfil a wide range of applications and provide for the monitoring and/or control of moisture in gases. The instrument consists of two component parts: monitor and sensor (TRANSMET I.S. transmitter), each of which is individually calibrated to a single standard allowing combinations of sensor (TRANSMET I.S. dewpoint transmitter) and monitor units to be totally interchangeable.

The standard instrument covers the ranges –100 (Cdp to +20 (Cdp (-148 (Fdp to +68 (Fdp), 0 to 1000 parts per million by volume (ppmV) as well as 0 to 1000 LB/MMSCF and g/m3 for natural gas. Selection of the displayed hygrometric unit is factory set but may be easily changed by the user. Dual alarm relay contacts are provided which are user configurable both in terms of setpoint and operating mode. Optionally, 1 or 2 further relays can be fitted with Form A type contacts (see section 2.2.2 for details). Current output is standard and factory set at 4 – 20 mA, optional outputs are 0 – 20 mA or

0 - 10 V DC.

The monitor is fitted with a pressure input channel to accept the signal from any loop-powered intrinsically safe 2-wire pressure transmitter. Pressure can be scaled and displayed to appropriate engineering units. The pressure signal can then be used to provide pressure compensation on the primary channel, of either LB/MMSCF, g/m3 or ppmV.

1.3	TRANSMET I.S. Overview

The TRANSMET I.S. Ceramic Dewpoint Transmitter measures dew-point temperature over  the range -100/+20 (Cdp. It is designed for use within hazardous areas: zones 0,1 & 2 when used with an appropriate safety barrier(s), and according to EN50 014 and EN50 020.

The TRANSMET I.S. is connected as a 3-wire transmitter where a current sink output can be switch selected for either 4-20 mA or 0-20 mA. This output is linear over the entire dew-point measurement range, i.e. -100/+20 (Cdp.

Note that the CERMET II monitor is set-up to accept a 4-20 mA input as standard.

TRANSMET I.S. is certified intrinsically safe for use in hazardous areas, by E.E.C.S. (BASEEFA). The instrument conforms to the C.E.N.E.L.E.C. standards specified in BS EN 50014 and BS EN 50020, with certification code EEx ia IIC T4.

Before using TRANSMET I.S. in any hazardous environment, ensure you are fully familiar with the above CENELEC standards relating to the certification of this instrument. Also be familiar with further information relating to the installation of intrinsically safe apparatus to be found in the BS EN 60079-14: 1997 or equivalent codes of practice in the country of installation.

TRANSMET I.S. utilises a Ceramic Moisture Sensor manufactured from metallised ceramics using thin and thick film technologies. The Michell ceramic moisture sensor is virtually chemically inert with inherently fast response, high calibration stability and high resilience to corrosive environments. TRANSMET I.S. measures humidity by monitoring the electrical response exhibited by the sensor to variations in partial pressure of water vapour of the gas composition to which it is exposed.

## 1.4	Outline of Complete System

Figure 1. Complete System

For connection details, refer to Figure 7.

1.5	Calibration

Michell Instruments’ calibration facilities are amongst the most sophisticated in the world and have been recognised for their excellence by the National Physical Laboratory (NPL) UK. Our humidity calibration laboratory is accredited by UKAS for dew-point measurements in the range –75 (Cdp to +20 (Cdp (calibration accreditation number 0179) and is traceable directly to the National Institute of Standards & Technologies (NIST) USA.

Furthermore, our systems provide calibration capability down to –100 (C dew point [13.8 ppb(V) at atmosphere pressure] with direct reference to a fundamental cooled mirror dew point hygrometer. Our high humidity facility provides calibration capability to +82 (Cdp @ 85 (C ambient.

1.6	Manufacturing Quality

Your instrument should reach you in perfect working condition. Michell Instruments Ltd are registered with the British Standards Institute for Quality Assurance to BS EN ISO 9002: 1994 – certificate number Q06284. We have rigorous procedures at every stage of production to ensure that the materials of construction, manufacturing, calibration and final test procedures meet the requirements laid down in our Quality Control Manual.

2.	PREPARING YOUR CERMET II MONITOR FOR USE

The CERMET II I.S. Hygrometer is supplied with a full accessory package. Please check that all the items indicated on the packing check list are present in the delivery.

NOTE:	The CERMET II Monitor is NOT in itself INTRINSICALLY SAFE and therefore must not be installed in the hazardous area.

2.1	Installing the Monitor

The monitor can operate within the environment limits of 0 (C to 50 (C, 0 to 90% RH.

The monitor case is designed for panel mounting. However, it can be used as a bench mounted device without any special preparation.

For panel mounting, a suitable cutout should be made in the panel to be used. A rectangular cutout of 92 x 45 mm (DIN 1/8) is necessary (see Figure 2. Panel Cut-out). Optional IP66/NEMA 4 cover provides additional front panel protection for harsh environments.

NOTE:	A minimum depth behind the panel of 140 mm is required for the monitor, its connectors and wiring (see Figure 3. Case Dimensions).

The monitor should be inserted into the panel cutout from the front and secured using the two brackets with securing screws on the sides of the front case.

All electrical connections through the rear panel are made via plug and socket terminals to facilitate easy removal of the monitor when in service.

Figure 2. Panel Cut-out

2.2	Electrical Connections

Figure 3. Case Dimensions

When the monitor is securely fastened in position, the electrical connections can be made to the back panel.

Figure 4. Instrument Wiring Diagram

Figure 4 above shows the connections for power, dew-point sensor, pressure sensor, alarm relays and analogue output. These details are also repeated on the instrument itself. The exact factory setup of the monitor unit can also be obtained by reference to the tick boxes shown on the connection details label, which is applied to the monitor case.

### 2.2.1	Electrical Power

The power supply to the monitor may be one of the following:

- 85-265 V AC 50/60 Hz and 95-370 V DC - factory default setting

- 18-36 V AC and 9-60 V DC - optional

The power supply voltage is indicated on the connection details label located on the monitor.

As soon as power is connected at the rear of the monitor, the display will register a reading. The monitor is provided for continuous operation and does not feature a power ON/OFF switch.

### 2.2.2	Setpoint (Alarm) Relays

SP1 and SP2 are Form C (single pole, change-over) relay contacts and can be connected as Normally Open or Normally Closed. Connections are NO - Normally Open, NC - Normally Closed and COM - common.  Contacts are rated 240 V AC: 10 A or 24 V DC: 8 A (non-inductive load).

Optional additional relays SP3 and SP4 have Form A (single pole, normally closed) relay contacts and are Normally Closed only. Connections are NC - normally closed and COM - common.  Contacts are rated 240 V AC: 5 A or 30 V DC: 5 A (non-inductive load).

Configuration of setpoints for operation as under \ over range, or band limit is covered in section 7.3.

### 2.2.3	Analogue Output

The monitor as standard provides a 4 to 20 milliampere (optionally: 0 to 20 mA or

0-10 V DC) linear output over the entire operating range of the instrument in the selected engineering units.  The output can be digitally scaled to cover a partial range of the measurement parameter. See CAL (under setup codes) in appendix 2.

The maximum load resistance that can be connected between pin 16(+) and pin 17(-) for a CURRENT output is 500 Ohms (No minimum).

The minimum load that can be connected to pin 16(+) and pin 17(-) for a VOLTAGE output is 5 Kilo Ohms (No maximum).

### 2.2.4	Digital Interface

The monitor is equipped with an RS232 digital interface as standard, or optionally an RS485 digital interface. A 2 metre RS232 or RS485 cable, terminated with a 9 way ‘D’ socket, is available on request.

3.	TRANSMET I.S.

3.1	TRANSMET I.S. Identification

The TRANSMET I.S. can be identified by a serial number label located on the outside of the TRANSMET I.S. housing. The label also describes the I.S. certification code for TRANSMET I.S.

## 3.2	Mechanical Installation

TRANSMET I.S. is provided with a 5/8"UNF parallel mounting thread which is suitable for housing in an optional flow-through sampling block. The body of TRANSMET I.S. is designed to mate with the bonded seal provided (optional stainless steel bonded viton seal available). With TRANSMET I.S. the bonded seal provided should be placed over the sensing part of the transmitter before it is screwed into the sampling block. TRANSMET I.S. can also be mounted by direct insertion. Remove the protective cap before installation and retain for future use.

If TRANSMET I.S. is to be mounted in the optional sampling block, appropriate gas couplings should be selected and fitted (using PTFE tape) into the female 1/8"NPT (taper) threads provided. Sample gas flow can be in either direction through the block. If the TRANSMET I.S. is to be mounted by direct insertion, appropriate care should be taken to ensure a proper seal.

Although the operation of TRANSMET I.S. is not sample flow-rate dependant it is important to ensure that the flow velocity through the sample line connecting the sample source to the sampling block is high enough to avoid long time lags in response to changes in humidity at the sample source. We therefore recommend a flow-rate of 1 to 5 litres/minute at sample tap-off point (or equivalent at pressure) be set when TRANSMET I.S. is mounted in the standard sampling block and that the instrument is mounted as close as practicably possible to the sample tap-off point. The gas sample may then be returned to the process gas stream via the 2nd port in the sampling block, or alternatively, if operating conditions allow, the gas sample can be vented - a vent pipe of at least 1 metre in length should be fitted to the outlet port of the sample block. In direct insertion applications a wide range of gas flow velocities are acceptable - Flow velocities up to 20 metres/second are acceptable with the protection given by  the standard sintered guard.

Note: The correct selection of cable gland for TRANSMET I.S. is essential with regards to I.S. installations. The specified IP66 rating of the instrument will only be maintained if an equivalent or higher IP rated cable gland is fitted. See TRANSMET I.S. specification for dimensions.

After installation into the gas stream the TRANSMET I.S. housing may be positioned at any angle through approximately 330( of rotation, to allow for the cable gland positioning. To position the Sensor housing first loosen the large clamping nut sufficient to allow free rotation of the housing around the Sensor body.

WARNING: TRANSMET I.S. is fitted internally with a stop device to prevent full 360( rotation of the housing around the Sensor body. If rotation is forced beyond 330( this may internally damage the Sensor. The stop device is only to provide indication to the installer when maximum travel in either direction has been reached.

Rotate the sensor housing until the cable gland is in the desired position. Whilst firmly holding the housing in position, retighten the large clamp nut up against the housing seal using a spanner of the correct size. Do not apply excessive force.

High pressure samples may be measured up to a maximum of 300 bar. Make sure that the correct pipe work fittings are used and that the bonded seal provided is used to seal the radial face of TRANSMET I.S./ sensor block seal.

Similarly TRANSMET I.S. is suitable for measuring samples below atmospheric pressure. Care should be taken at joints and pipe fittings that no leak exists which could allow ambient air to contaminate the sample gas. If there is doubt a standard leak test procedure should be employed (helium or vacuum seal).

## 3.3	Operating Precautions

Precautions should be taken to protect the Ceramic Moisture Sensor from damage. Statistical information indicates that the vast majority of failures are caused either by incorrect sampling methods, sampling positions or inadequate external protection against dangerous substances and aggressive gas borne contaminants.

## 3.4	Sampling Hints

TRANSMET I.S. may be provided with a sampling block that will enable a small sample of process gas to be diverted past the ceramic sensing element before returning to the main gas stream or being bled off to atmosphere. Alternatively TRANSMET I.S. can be mounted directly into a flowing gas stream in a duct or pipe, provided there is not any form of contamination within the duct/pipe that will damage the ceramic sensing element.

Sample system guidelines:

a)	Ensure the sample is representative of the gas under test.

The sample point should be as close to the critical measurement point as possible.

b)	Minimise dead space in sample lines.

Minimise couplings and pipe work. Where possible, build-up sampling pipe specifically for the job and do not use pipe work previously installed for another application. Dead space in sample lines increases response time by holding water molecules that are released to the passing gas sample.

c)	Remove any particulate matter or oil from the gas sample.

The Ceramic Moisture Sensor element of TRANSMET I.S. is protected against particulate contamination by a 80(m st. st. sintered guard. This protects against physical damage caused by large particles at high velocity such as rust in a compressed air line. However, fine particles will not be prevented from contact with the Ceramic Sensing Element. If high concentrations of fine particles are present in the sample flow they may “blind” the Ceramic Moisture Sensor element and reduce its response speed. If TRANSMET I.S. is installed in a system where particulate such as degraded desiccant or pipe swarf and rust are present, use a particulate in-line filter. Contact Michell Technical Sales staff for further advice on filtering techniques.

d)	Use high quality sample pipe and fittings.

We would recommend that, wherever possible, stainless steel pipe work, fittings and components (e.g. filters) should be used. This is particularly important when measuring low dew points since materials other than st. st. may have undesirable characteristics such as the absorbing of moisture on the walls, which slows down response and in extreme circumstances gives false readings. For temporary applications, or where stainless steel pipe work is not feasible, use high quality thick-walled p.t.f.e. pipe work as this exhibits similar characteristics to stainless steel.

Use the lowest internal diameter pipe work possible to reduce response time, but take care not to induce pressure differentials by aiming for too high a flow rate through small bore pipe work. A sampling flow up to 1 to 5 L/min (or equivalent at pressure) will be satisfactory for TRANSMET I.S. to operate correctly.

3.5	Response Characteristics

Response characteristics from dry to wet are orders of magnitude faster than from wet to dry.

Time taken to dry down TRANSMET I.S. from ambient conditions to the operational dew-point level of the process will normally be shorter than the time taken to dry down the process itself. Therefore when TRANSMET I.S. is installed into the system prior to system start-up there is normally no time lag before representative test results are obtained.

When a TRANSMET I.S. is installed into an operational system, then typically fifteen to thirty minutes should be allowed for any pipe work, filter and TRANSMET I.S. to reach equilibrium with the sample gas passing through.

3.6	Which Gases to Measure

TRANSMET I.S., by nature of its design, is suitable for measurement of the moisture content of a wide variety of gases. In general, if the gas (in conjunction with water vapour) is not corrosive to base metals then it will be suitable for measurement by the Transmitter. However gases containing entrained solids or hydrocarbon mists should be filtered before presentation to the sensor using a coalescing filter. If only particulate matter is present in the sample line then a particulate in-line filter is all that is required. Suitable filters are available from Michell Instruments or from a Michell approved Sales and Service Centre.

If a very dry gas, which would become corrosive when coming into contact with higher concentrations of moisture, is to be measured, then the instrument must be purged with an inert gas immediately after the measurement has been made. This will prevent any potentially corrosive gases remaining in the sensor block from contaminating the sensor.

3.7	Maintaining the TRANSMET I.S.

Routine maintenance of TRANSMET I.S. should only require regular re-calibrations. The normal calibration period is one year. This re-calibration work can only be performed by exposure of the moisture sensor to sample gases of known moisture content. Calibration services traceable to national humidity standards are provided by Michell Instruments' UKAS accredited calibration laboratory located in Cambridge and at our sales and service centres located in Denmark, Holland, Germany, Japan and USA - contact Michell Instruments or your local representative for further details.

The frequency of re-calibrations required in order to maintain the performance of TRANSMET I.S. transmitter is primarily dependant on the composition of the gas to which the Moisture Sensor is exposed, i.e. content of liquid and particulate contaminates, corrosive elements, etc. (refer to section 3 of this manual for guidance). In most applications annual re-calibration ensures that the stated accuracy of TRANSMET I.S. is maintained. Michell Instruments offer a comprehensive package of Service Contracts which can be tailored to customers specific needs.

4.	TRANSMET I.S. TERMINAL BLOCK CONNECTION

Connections to the TRANSMET I.S. are made via 3 terminals of an 8-way screw terminal block located on the TRANSMET I.S. PCB (annotated TB1) as shown below in figure 5 below. The terminal block has its terminals labelled 1 to 8, however, only terminals 5,6 and 8 are required to connect the TRANSMET I.S. To gain access to the terminal block, unscrew the lid of the TRANSMET I.S housing.

Figure 5. Terminal Block Connections

5.	TRANSMET I.S. CURRENT OUTPUT

There are two current ranges available from the TRANSMET I.S. which are switch selectable, as either 4 to 20 mA or 0 to 20 mA, both corresponding to a dew point of
-100 to +20 (Cdp. An 8-way DIL switch is located on the TRANSMET I.S. PCB and is accessed by removing the TRANSMET I.S. housing lid. The 8-way DIL switch is annotated SW1. (See Figure 5 below). Switches 1 to 5, 7 & 8 on the 8-way DIL switch must be off. Switch 6 is used to select the current range. By default, the CERMET II monitor is setup to accept a 4-20 mA range input.

Note: Current range can be changed whilst the TRANSMET I.S. is in operation.

Figure 6. Current Output Switch Settings

6.	CONNECTING TRANSMET I.S. TO CERMET II MONITOR

SAFETY FIRST

TRANSMET I.S. is Certified intrinsically safe for use in Hazardous areas, by E.E.C.S. (BASEEFA). The transmitter conforms to the C.E.N.E.L.E.C. standards specified in BS EN 50 014 and BS EN 50 020, with certification code EEx ia IIC T4.

Before using TRANSMET I.S. in any hazardous environment, ensure you are fully familiar with the above C.E.N.E.L.E.C. standards relating to the certification of this instrument and also with the further information relating to intrinsically safe apparatus to be found in the BS5345 Code of Practice, parts 1 & 4 or equivalent codes of practice in the country of installation.

NOTE: Installation of TRANSMET I.S. MUST be as per system 2 drawing in order to comply with the Intrinsic safety certification for TRANSMET I.S.

Figure 7. Connecting TRANSMET I.S. to CERMET II Monitor

Figure 7 above shows a complete system installation including an intrinsically safe pressure transmitter.

7.	CERMET II MONITOR SETUP

There are two levels of operation:

- User, which does not allow access to the programming codes;

- Advanced, which allows access to the programming codes.

User allows:

- Changing of the display brightness;

- Monitoring of the setpoint values.

Advanced programming allows setting of the following options:

- advanced setpoint programming;

- analogue output calibration and scaling;

- auxiliary (optional) channel setup;

- setting the engineering units;

- other related advanced functions.

## 7.1	Setup Security Feature

To prevent unauthorised access to the internal settings, the monitor is equipped with two DIP-switches located on the display board of the monitor, which can be accessed underneath of the faceplate (refer to Figure 8. Location of the LOCKOUT Switches.). The PROGRAM LOCKOUT switch SW2 is responsible for enabling or disabling the code programming mode. The SETPOINT LOCKOUT switch SW1 enables or disables the alarm setpoint programming. OFF position of either switch will enable that mode, ON position will disable that mode.

Figure 8. Location of the LOCKOUT Switches.

## 7.2	Selecting the Engineering Units

Enter the PROGRAM UNLOCK mode as described in 7.1	Setup Security Feature, and press SETUP and ( to scroll through the menus.

The monitor can display dew point in (C or (F, PPM(V), LB/MMSCF or g/m3 (Natural Gas). The selection of these units is achieved by setting CODE 4 and CODE 7 as shown below: (See Appendix 2 – CODE 4 & CODE 7 for full details).

To display Dew Point set CODE 7 		to 	000

To display Dew Point in (C set CODE 4 		to 	207

To display Dew Point in (F set CODE 4 		to 	217

To display PPM(V)  set CODE 7 			to 	100

To display LB/MMSCF  set CODE 7 		to 	200

To display g/m3  set CODE 7 			to 	300

To set the range and resolution for PPM(V), LB/MMSCF and g/m3 , set digit 3 of CODE 7 to :-

0 	1 – 9999

1	0.1 – 999.9

2	0.01 – 99.99

3	0.001 – 9.999

For example, by setting CODE 7 to 102, displays PPM(V) with a resolution of 0.01 to a maximum of 99.99, above which, oVEr will be displayed.

## 7.3	Changing the Setpoint Values

Before amending the setpoint values, make sure that the monitor is in the SETPOINT UNLOCK mode, as described in 7.1 Setup Security Feature.

Press SETUP and ( simultaneously to gain access to the setpoint codes.

The setpoints of the relays are set by setting SP_n (where n = 1 to 4 and represent SP1 to 4), to the required level. The units of the setpoints, change according to the source of the setpoints. i.e If the indicator is displaying ppmv, then the setpoints are set in ppmv.

For the relays to be energized above the setpoint  SPC_n = 0xx.

For the relays to be energized below the setpoint  SPC_n = 1xx.

The source of relay setpoints are set by setting digit 2 of SPC_n  i.e xnx as shown below:

0	Displayed value  (DEFAULT)

1	Dew Point

2	Pressure

3	Not Applicable for Cermet II I.S.

- Not Applicable for Cermet II I.S.

- Not Applicable for Cermet II I.S.

- Not Applicable for Cermet II I.S.

For example; if you wish the relay to trip at a certain dew-point value whilst the indicator is displaying ppmV, then SPC_n = x1x.

The secondary relay function is set by digit 3 of SPC_n  i.e xxn as shown below:-

0	No function   (DEFAULT)

1	Relay latched

2	De-energize relay

3	No function

4	Not Applicable for Cermet II I.S.

5	Not Applicable for Cermet II I.S.

6	Not Applicable for Cermet II I.S.

7	Setup for hysteresis, delay type and make \ break delay times

For example, if you wish the relay to trip above the setpoint using the displayed value, but to remain latched, then SPC_n = 001. It will then remain latched until the supply to the indicator is removed or SPC_n = 002.

The function of the relay annunciators is set by digit 1 of CODE 1 i.e nxx. As shown below:

0	LED annunciators  always OFF

1	LED annunciators ON when relays are de-energized

2	LED annunciators ON when relays are energized (DEFAULT)

Hysteresis, Make\Break delay & delay type

Associated with each setpoint is a Hysteresis Value, Make delay time, Break delay time and a delay type. To gain access to these parameters, set SPC_n  = xx7, and to scroll the features press ( or (.

The hysteresis value is a value which the relay trips above and below the nominal setpoint. For example, if SP1 is set to -10.0 and the HYST is set to 2.0 and the relay is set to energize above the setpoint, then the relay will energize when the setpoint source rises above -8.0 and is de-energized when the setpoint source falls below

-12.0

The Make delay is the delay between the setpoint being reached and the relay energizing. For example, if M_d = 0.02.30 then the relay will be energized 2 minutes and thirty seconds after the trip point has been reached.

The Break delay is the delay between the relay being energized and it de-energizing. For example, if b_d = 0.01.12 then  when the relay is energized it will remain energized for 1 minute and 12 seconds , after which it will be de-energized.

The maximum make and break delay time is 9 hours 6 mins and 6 secs in increments of 1 second.

The setpoint can have four different delay types:- NorM (normal), rEPt (repeat), 1Shot, and PuLSE.

If deLAY = NorM then the relay will function normally with the inclusion of the time delays.

If dELAY = rEPt then the make and break delays will repeat continually until the setpoint source goes to a level to deactivate the setpoint. For example, if the make delay = 10 secs & break delay = 5 secs, then after the trip point is reached, the relay will be energized after 10 secs and remain energized for 5 secs after which it will de-energize for 10 secs and then energize for 5 secs etc.

If dELAY = PuLSE then if using the make and break times mentioned above, the relay will energized after 10 secs, de-energize after 5 secs and remain de-energized.

If dELAY = 1Shot then if using the make and break times mentioned above, the relay will energize after 10 secs and remain energized.

7.4	Analogue Output Scaling

Enter the PROGRAM UNLOCK mode as described in 3.1, press SETUP then (, then SETUP again. Now use ( and ( to scroll through to the CAL menu.

The analogue output can be scaled by setting CAL to 061 and setting the zero value to the required output at 4 mA ( or 0 V) and the full scale value to the output required at

20 mA ( or 10 V).

For example, if the output is required to go from –90.0 (Cdp at 4 mA to +10 (Cdp at

20 mA then set Zero to –90.0 and F.S. to 10.0

To exit the CAL menu press SETUP, set CAL to 000 and use ( to scroll through the menu’s.

In order for the analogue output to represent the value displayed by the indicator, i.e. Dew Point, PPM(V), LB/MMSCF or g/m3, then set CODE 3 to 000. The values for Zero and F.S. are then set in the appropriate units i.e. in PPM(V) when PPM(V) is the value being displayed.

If the analogue output is required to represent dew-point whilst the unit is displaying one of the calculated values (PPM(V), LB/MMSCF or g/m3) then set CODE 3 to 010.

If the analogue output is required to represent pressure, then set CODE 3 to 020.

## 7.5	Display Brightness Adjustment

To adjust the display brightness:

-press SETUP and ( buttons simultaneously, the display toggles between [bri] and [5], where 5 is default setting;

-adjust the display brightness required (from 0 to 7) by using ( or ( buttons;

-press SETUP several times to exit the programming mode.

## 7.6 Digital Communications

The monitor is fitted with an ASCII RS232 as standard or an optional RS485.

The communication settings can be found by setting CAL to 100, then, press the SETUP button once and then ( and ( to change the baud rate. Baud rates are 300, 600, 1200, 2400, 4800, 9600, 19,200 and 57.6 K. Press the SETUP button again to advance to the parity bit and then ( and ( to change the parity bit. Parity settings are odd, even or off. Press the SETUP button again to advance to the address settings. Address settings are 0 to 255. Note, address 0 is not a valid RS485 address. Press the SETUP button again to return to CAL.

The communications uses the following default protocol:

Baudrate 	9600

Data bits	8

Parity		none

Flow control	Xon\Xoff

The message to read and write to the monitor follows the following format:

1) Start Character

"s" or "S" for the start character (must be first character in string).

2) Meter Address

An ASCII number from "0" to "255" for the meter address. If the character following the start character is not an ASCII number, then address 0 is assumed. All meters respond to address 0.

3) Read/Write Command

The next character must be an ASCII "R" or "r" for read, or an ASCII "W" or "w" for write. Any other character will abort the operation.

4) Register Address

The register address for the read/write operation is specified next. It can be either an ASCII number from "0" to "255" or registers 1 - 18 can be accessed by entering and ASCII letter from "A" to "R" (or "a" to "r", not case sensitive). If the address character is omitted in a read command, the meter will always respond with the data value currently on the display. (The register address must be specified for a write command).

5) Separator Character

After the register address in a write command, the next character must be something other than an ASCII number. This is used to separate the register address from the data value. It can be a space or a "," or any other character except a "$" or a "*".

6) Data Value

After the separator character, the data value is sent . It must be an ASCII number in the range of "-32766" to "32766".

7) Message Terminator

The last character in the message is the message terminator and this must be either a "$" or a "*". If the "$" is used as a terminator, a minimum delay of

50 mS is inserted before a reply is sent. If the "*" is used as a terminator, a minimum delay of 2 mS is inserted before a reply is sent. (the
"$" and "*" characters must not appear anywhere else in the message string).

Examples:

SR$			Read display value, 50 mS delay, all meters respond.

s15r$			Read display value, 50 mS delay, meter address 15 responds.

Sr130*		Read Code 1 setting, 2 mS delay, all meters respond.

s2wb-10000$	Write -10000 to the display register of meter address 2, 50 mS delay.

SW65 200$		Write 200 to the SP1 Hysteresis register, 50 mS delay, all meters respond.

s10w148,7*		Change brightness to 7 on meter address 10, 2 mS delay.

sr6$			Reads SP1

sw6-100$		Sets SP1 to –10.0

sr5$			Reads pressure value

A full list of the register that can be read or written to can be found in Appendix 1.

8.	PRESSURE COMPENSATION

Enter the PROGRAM UNLOCK mode as described in section 7.1.

The monitor has the ability to measure pressure in order to provide a pressure compensated value for PPM(V), LB/MMSCF or g/m3.

## 8.1	Using a Pressure Transducer

In order to enable the use of a pressure transducer set CODE 4 to 300 (dual channel mode), by pressing SETUP and ( to scroll through the menu’s.

If you wish to display Dew-point whilst measuring pressure, then set CODE 4 to 307 to show Dew Point in (C or 317 in (F.

To display pressure in PSIG set CODE 6 to 000., or set CODE 6 to 100 to display in BARG.

Manual pressure input calibration

The pressure input channel needs to be calibrated in order to suit the range of the pressure transducer. This is achieved by setting CAL to 012 and entering values for offset OFF_2 and scale SCA_2.

SCA_2 = 0.0062 per 100 PSIG

OFF_2 = pressure range – ((20,000 x (pressure range\1000))\16)

For example, for a pressure transducer with a range of 0 to 1000 PSIG

SCA_2 = 0.0062 x 10

= 0.0620

OFF_2 = 1000 – ((20,000 x (1000\1000))\16)

= 1000 – 1250

= -250

If you are using a pressure transducer in BARG, then you have to convert the value to PSIG by multiplying it by 14.5.

For a pressure transducer with a range of 0 to 100 BARG :-

100 BARG = 1450 PSIG

SCA_2 = 0.0062 x 14.5

= 0.0899

OFF_2 = 1450 – ((20,000 x (1450\1000))\16

= 1450 – 1812

= -362

To exit the CAL menu, set CAL to 000 and press ( to scroll through the menus.

Automatic pressure input channel

The pressure input can alternatively be calibrated using a 4 and 20 mA DC current source. To do this, connect the current source between pin 3 (-ve) and pin 2 (+ve).

Set CAL to 022 and press SETUP. The monitor will then display Zero and flash 0. Set the mA source to 4.0 mA and press SETUP to set 0 at 4 mA.

Press SETUP and the monitor will display SPAn and flash the full scale pressure value. Set the mA source to 20.0 mA and use the ( & ( buttons to set the required pressure range. The press SETUP to set the required pressure for 20.0 mA.

Set CAL to 000 and press ( to exit the menus.

## 8.2	Using a Fixed Pressure Input in Single Channel Mode

In order to display pressure compensated values of PPM(V), LB/MMSCF and g/m3 without the use of a pressure transducer, then this can be achieved by entering the pressure value manually.

Enter the PROGRAM UNLOCK mode. Use SETUP and ( to scroll through the menus and select code 7.

Select the required pressure compensated value, i.e. CODE 7 = 1xx for ppmV, 2xx for lbmmscf or 3xx for g/m3 , and CAL = 052.

Exit the menus and exit the PROGRAM UNLOCK MODE. Press SETUP for a two seconds, and the display will show the set pressure. To change the pressure, use ( and ( and then press SETUP.

## 8.3	Pressure Transducer Connection

The monitor provides excitation voltage (24 V DC @ 20 mA) for an auxiliary 2-wire transmitter (via an intrinsically safe barrier unit) used in the dual channel configuration. Connect (+) of the intrinsically safe barrier to pin 2 of the monitor, (-) of the intrinsically safe barrier to pin 3 of the monitor. Refer to figure 7 for details.

9.	MONITOR MAINTENANCE

The CERMET II I.S. Hygrometer does not require regular re-calibration or servicing.
However, if a fault with the monitor is suspected, contact Michell Instruments technical support.

10.	TECHNICAL SPECIFICATIONS

## 10.1	CERMET II Monitor

Display:				5 digit LED

Dew-point Measurement			-100 (Cdp to +20 (Cdp (-148 (Fdp to +68 (Fdp)

Ranges:				0 to 9999 PPMV, 0 to 1000

LB/MMSCF, g/m3 (Natural Gas).

Pressure input channel:		4-20 mA input scaleable to any engineering units.

Electronic accuracy:		Dew-point ranges ±0.5(C, ±1.0(F, ppmV range ±1% of reading, LB/MMSCF range ±1% of reading, secondary channel ±0.5% FS.

Outputs:				4-20 mA as standard or 0-20 mA optional (max load 500 ohms) or 0-10 V optional (min load 5 K ohms), spanned to full calibration range, scaleable by user.

Alarm Relays:			Two relays SP1 and SP2 are standard. Control action and setpoint are user programmable.  Form C contacts rated 10 A, 240 V AC or 8 A, 24 V DC. Non-inductive load.

Relays SP3 and SP4 are optional. Control action and setpoint are user programmable.  Form A contacts rated 5 A, 240 V AC or 5 A, 30 V DC. Non-inductive load.

Operating environment:		0( to +50 (C, 0 to 90% RH.

Power supply:			82-265 V AC, 50/60 Hz or 95-370 V DC standard.

Optionally: 18-36 V AC or 9-60 V DC.

Power consumption:		Max 10 Watts.

Power connection:			2 metre, 3 wire.

Weight:				0.6 Kg.

Environmental protection:	IP54 / NEMA 12. Optional protection cover to IP66 / NEMA 4.

## 10.2	TRANSMET I.S.

| I.S. Certification code | EEx ia IIC T4. To B.S. EN 500 14 and B.S. EN 500 20 |
| --- | --- |
| Sensor type | Michell Ceramic Moisture Sensor |
| Gas wetted components | 316 Stainless steel |
| Sensor mounting thread | 5/8”UNF |
| Housing | 316 Stainless steel natural finish |
| Ingress protection | IP66, NEMA 4 |
| Electrical connection | Internal screw terminals. Connection Capacity: Solid / Stranded / Conductor sizes: mm2/mm2/AWG = 0.14 - 0.5 / 0.14 - 0.5 / 26 - 20 |
| Dimensions | Outline: 172mm long x 106mm high x 86mm diameter (lid) |
| Weight | 1.1 Kg |
| Cable gland port | ½” BSP * |
| Power | Max. 30 V regulated DC supply Min. 6 V regulated DC supply Current consumption 15 mA ( 30% (not including dew-point signal current sink) |
| Operating modes | Selectable 0-20 mA or 4-20 mA linear current sink output Corresponding to (C dew point |
| Range | -100 / +20 (Cdp (-148 / +68 (Fdp) |
| Accuracy | (1 (C between +20 (Cdp & -60 (Cdp (2 (C between –60 (Cdp & -100 (Cdp |
| Resolution | 0.1 (C between +20 (Cdp & -80 (Cdp 1 (C between –80 (Cdp & -100 (Cdp |
| Operating temperature | -20 / +40 (C (-4 / +104 (F) |
| Storage temperature | -40 / +70 (C (-40 / 158 (F) |
| Operating pressure | max 300 bar g |
| Flow rate | 0 to 10 NL/min-1 |
| Gas Velocity | max 20 ms-1 |

* Note that thread adapters to convert the cable gland port to either ¼”NPT or M20 x 1.5 are available on request from M.I.L.

11.	INTRINSICALLY SAFE SYSTEM 2 DRAWING

Isolation barrier installation method

APPENDIX 1

### Register Settings Accessible by Digital Communication

Register number		Function					Read only

- Alarm Status

- N/A

- Processed Data – Result 				(

- Processed Data – Channel 1 (dew-point)	(

- Processed Data – Channel 2 (pressure)		(

- Setpoint 1

- Setpoint 2

- Setpoint 3

- Setpoint 4

10 to 23 			N/A

- Scale Value – Result

- Scale Value – Channel 1

- Scale Value – Channel 2

- Offset Value – Result

- Offset Value – Channel 1

- Offset Value – Channel 2

30 to 33			N/A

- D/A Zero – Analogue O/P 1

- N/A

- D/A Full Scale – Analogue O/P 1

37 to 64			N/A

- Hysteresis - Setpoint 1

- Hysteresis - Setpoint 2

- Hysteresis - Setpoint 3

- Hysteresis - Setpoint 4

69 to 70			N/A

- Make Delay - Setpoint 1

- Make Delay - Setpoint 2

- Make Delay - Setpoint 3

- Make Delay - Setpoint 4

75 to 76			N/A

- Break Delay - Setpoint 1

- Break Delay - Setpoint 2

- Break Delay - Setpoint 3

- Break Delay - Setpoint 4

81 to 128			N/A

- Cal Mode

- Code 1

- Code 2

- Code 3

- Code 4

- Code 5

- Code 6

### Register Settings Accessible by Digital Communication cont.

Register number		Function				Read only

136				Code 7

- Code 8

- to 141		Reserved

- Setpoint 1 Control Register

- Setpoint 2 Control Register

- Setpoint 3 Control Register

- Setpoint 4 Control Register

146 to 147			N/A

- Brightness

- Baudrate Settings

- Serial Address

151 to 152			N/A

153				Model Number				(

154				Version Number				(

155 to 192			N/A

- Delay Type – Setpoint 1

- Delay Type – Setpoint 2

- Delay Type – Setpoint 3

- Delay Type – Setpoint 4

APPENDIX 2

### SETUP Codes

CAL

Calibration modes for input and output.

1st digit (left most):	Calibration mode

Digit:		Function

- Calibration functions as per 2nd and 3rd digit

- Set baud rate, parity and serial address

2nd digit:		Calibration function

Digit:		Function

- No function

- Manual calibration (channel as per 3rd digit)

2 to 4		N/A

5		Manual adjust

- Analogue output scaling (analogue channel as per 3rd digit)

- N/A

3rd digit:		Object for calibration

Digit:		Function

- No function

- Processed result (dew point, ppmV, lb/mmscf, g/m3)

- Pressure input

DEFAULT VALUE = 052

CODE 1

Tendancy Indication, Additional LED’s, Display data source, Display flashing, Decimal points, Display rounding.

1st digit (left most):	additional LED’s

Digit:		Function

- LED annunciators are always off

- LED annunciators are on when relays are de-energized

- LED annunciators are on when relays are energized

- N/A

2nd digit:		Display data source

Digit:		Function

- Processed data - result (dew point, ppmV, lb/mmscf, g/m3)

- Processed data - channel 1 (dew point only)

- Processed data - pressure

3 to7		N/A

3rd digit:		0 only

DEFAULT VALUE = 200

CODE 3

Serial mode and analogue output source

1st digit (left most):	Serial mode

Digit:		Function

- Ascii mode

1 to 3		N/A

2nd digit:		Analogue output 1 source

Digit:		Function

0		Analogue output 1 from processed result data (dew point, ppmV, lb/mmscf, g/m3)

1		Analogue output 1 from processed data (dew point only)

2		Analogue output 1 from processed data (pressure)

3 to 7		N/A

3rd digit:		0 only

DEFAULT VALUE = 000

CODE 4 – Channel 1

Measurement task, sampling rate

1st digit (left most):	Analogue sample rate

Digit:		Function

0 to 1		N/A

2		Single channel (50 Hz)

3		Dual channel (50 Hz)

2nd digit:		Analogue output 1 source

Digit:		Function

- Dew point in degrees C

- Dew point in degrees F

3rd digit:		0 only

DEFAULT VALUE = 207

CODE 6 – Channel 2

Measurement task

1st digit (left most):	Measurement task

Digit:		Function

0		Pressure in PSIG

1		Pressure in BARG

2		N/A

2nd digit:		0 only

3rd digit:		0 only

DEFAULT VALUE = 000

CODE 7

Result processing

1st digit (left most):	Measurement task

Digit:		Function

0		Dew point

1		PPMV

2		LB/MMSCF (natural gas)

3		g/m3 (natural gas)

2nd digit:		0 only

3rd digit:		Range for result

Digit:		Function

- 1 – 9999

- 0.1 – 999.9

- 0.01 – 99.99

- 0.001 – 9.999

DEFAULT VALUE = 000

Setpoint control 1 – 6 (accessed from Prog/Down menu after setpoints)

Relay latching, relay setup, source

1st digit (left most):	Relay sense

Digit:		Function

0		Relay energized above setpoint value

1		Relay energized below setpoint value

2 to 3		N/A

2nd digit:		Setpoint source

Digit:		Function

0		Processed result data (dew point, ppmV, lb/mmscf, g/m3)

1		Dew Point only

- Pressure

- Not Applicable for Cermet II I.S.

- Not Applicable for Cermet II I.S.

- Not Applicable for Cermet II I.S.

- Not Applicable for Cermet II I.S.

3rd digit:		Setpoint function for SP1 – SP4

Digit:		Function

- No function

- Relay latched

- De-energized relay

- N/A

- Not Applicable for Cermet II I.S.

- Not Applicable for Cermet II I.S.

- Not Applicable for Cermet II I.S.

- Setup hysterisis, make/break delay and delay type

DEFAULT VALUE = 000

APPENDIX 3

RECYCLING

Michell Instruments Limited is concerned with the protection of the environment. It is our commitment to reduce and eliminate from our operations, wherever possible, the use of substances which may be harmful to the environment. Similarly, we are increasingly using recyclable and/or recycled material in our business and products wherever it is practical to do so. The product you have purchased may contain recyclable and/or recycled parts and we will be happy to provide you with information on these components should you desire it.

## APPENDIX 4

## Hazardous Products

The Consumer Protection Act 1987, Section 6 of the Health and Safety at Work Act 1974 and the Control of Substances Hazardous to Health Regulations 1988, require that we advise the recipients and users of our products of any potential hazards associated with their storage, handling or use.

As the product detailed in this instruction manual contains a large number of components, both metallic and non-metallic, it is not practicable to issue full details of either all their constituents, substances or of the circumstances in which any of them may become hazardous.

We can, however, state that the product detailed in this manual and all of our other products are not hazardous to health when stored and used within the technical and environmental limitations specified in our relevant catalogue or specification sheet.

Should you require any further specific information regarding individual components of this product, please contact our Technical Sales Department.

## APPENDIX 5

## Certification

Certificate of Conformity & Schedule to BAS No. Ex 98E2129X (3 pages) dated

13 July 1998 and Variation to BAS No. Ex 98E2129X/1 (2 pages) dated 21 October 1998.

Certificate of Conformity & Schedule to BAS No. Ex 98E2144 (3 pages) dated

13 July 1998, Variation to BAS No. Ex 98E2144/1 (2 pages) dated 25 May 1999 and Variation to BAS No. Ex 98E2144/2 (2 pages) dated 11 July 2000.

EC-Type Examination Certificate No. BAS01ATEX1240X (3 pages)

dated 14 August 2001.

# Text boxes

1 2 3 4 5 6 7 8

TB1

SW1

TRANSMET I.S. PCB

Pin 8 = Current o/p

Pin 6 = +Power

Pin 5 = 0 V (GND)

1 2 3 4 5 6 7 8

TB1

SW1

TRANSMET I.S. PCB

ON =  0-20 mA output range

OFF

SW1

1

2

3

4

5

6

7

8

ON

OFF =  4-20 mA output range

- 11

2     12

I out

GND

8

5

1       7

2       8

6

+ PWR

TRANSMET I.S.

Ex1.50(p) barrier

Ex1.48 barrier

Hazardous area

Safe area

DP input

Pressure input

GND

1

2

3

4-20mA

Remote power supply

2

1

CERMET II monitor

+VE : 24V

-VE :GND

x       x

x       x

Barrier

4-20mA

I.S. Pressure transmitter

270R

1W

1%

270R

1W

1%

PSU

CERMET II

TRANSMET I.S.

SAFETY BARRIERS

I.S. PRESURE TRANSMITTER

SAFETY BARRIER

#### HAZARDOUS

#### SAFE